## Slide 1: Biyopotansiyel Elektrotlar
Yalova Üniversitesi
22 Kasım 2013

## Slide 2: Yarım-hücre gerilimi
Yarım-hücre gerilimi = Ehc = Ve
Elekrot-elektrolit arayüzeyinde oluşan iyon dengesizliklerinden dolayı
Üç ayrı tip madde üç şekilde birbiriyle temas haline getirilebilir
Katı–katı; ısıl çift (thermocouple)
Katı-sıvı; elektrot
Sııvı-sıvı; iyon seçici bir membranla ayrılmış, hücre gerilimlerinde olduğu gibi
Notes: Carr 6-11

## Slide 3: Ortak bir elektrolitik çözeltideki farklı iki metal
Elektrotlar arasındaki net gerilim
= VEA – VEB
Notes: Carr 6-12

## Slide 4: Elektrot-elektrolit arayüzü
Notes: W5-1

## Slide 5: Yaygın kullanılan elemanların yarım-hücre gerilimleri

## Slide 6: Bir gümüş-gümüş klorür elektrotunun kesiti

## Slide 7: Sinterlenmiş Ag/AgCl elektrot

## Slide 8: Polarizasyon
Overpotential = observed half-cell potential – equilibrium zero current half-cell potential
Vp = Vr + Vc + Va
Vp = total overpotential or polarization potential
Vr = ohmic overpotential
Vc = concentration overpotential
Va = activation overpotential

## Slide 9: Polarize olan ve olmayan elektrotlar
Perfectly polarizable
  No actual charge crosses electrode-electrolyte interface when current is applied
  The current is the displacement current as in a capacitor
Perfectly nonpolarizable
  Current passes through electrode-electrolyte interface freely without requiring additional energy
  No overpotentials
Neither is practically possible
  Electrodes made up noble metals such as platinum behave closest to perfectly polarizable
  silver-silver chloride electrodes behaves closest to perfectly nonpolarizable

## Slide 10: Gümüş-gümüş klorür elektrotların davranışı
Notes: Carr 6-13

## Slide 11: Elektrolitle temas halindeki bir biyopotensiyel elektrotun eşdeğer devresi
Notes: W5-4

## Slide 12: Elektrot-deri arayüzü ve hareket hataları
Derinin çeşitli kademelerinin büyütülmüş bir görüntüsü
Notes: W5-7

## Slide 13: Deri arayüzlü bir model
Sweat glands
and ducts
Electrode
Epidermis
Dermis and
subcutaneous layer
Ru
Re
Ese
Ehe
Rs
Rd
Cd
EP
RP
CP
Ce
Gel
Notes: W5-8

## Slide 14: Ag/AgCl elektrotların empedansının frekansla değişimi

## Slide 15: Elektrot empedansının frekansla değişiminin deneysel sonucu

## Slide 16: Biyopotensiyel kaydı için deri üzerine yerleştirilen iki biyopotansiyel elektrot
Notes: Carr 6-14a

## Slide 17: Bir biyomedikal kayıtın devre modeli

## Slide 18: Elektrot tipleri
Vücut yüzeyinden kayıt için elektrotlar
  Metal-plaka
  Emme (Suction)
  Yüzer (gömme – sütun) (Floating (recessed -column))
  Esnek (elastiki)
İç (internal) elektrotlar
  İğne Needle
  Tel
  Yerleştirilmiş (indwelling)
Elektrot dizileri
Mikro elektrotlar
  Metal
  Destekli metal
  mikropipet

## Slide 19: Biyopotansiyel vücut yüzeyi elektrotları
Metal plaka
Metal disk
EKG kayıtlrında sık kullanılan köpük arkalıklı tek kullanımlık elektrode
Notes: W5-9, Carr 6-15a

## Slide 20: Metalik emme elektrot
Genellilkle klinik elektrokardiograf uygulamalarında gögüs üzerine yerleştirilen bir metalik emme elektrodu
Notes: W5-10, Carr 6-

## Slide 21: Yüzer (gömme – sütun) elektrot
Recessed electrode with top-hat structure
Disposable recessed electrode
Double-sided
Adhesive-tape
ring
Insulating
package
Metal disk
Electrolyte gel
in recess
(a)
(b)
(c)
Snap coated with Ag-AgCl
External snap
Plastic cup
Tack
Plastic disk
Foam pad
Capillary loops
Dead cellular material
Germinating layer
Gel-coated sponge
Notes: W5-11

## Slide 22: Esnek yüzey elektrotları
Carbon-filled silicon rubber electrode
Flexible thin-film neonatal electrode
Cross-sectional view of thin-film electrode
Notes: W5-12

## Slide 23: Figure 5.13 Needle and wire electrodes for percutaneous measurement of biopotentials (a) Insulated needle electrode. (b) Coaxial needle electrode. (c) Bipolar coaxial electrode. (d) Fine-wire electrode connected to hypodermic needle, before being inserted. (e) Cross-sectional view of skin and muscle, showing coiled fine-wire electrode in place.

## Slide 24: FECG Electrodes
Figure 5.14 Electrodes for detecting fetal electrocardiogram during labor, by means of intracutaneous needles (a) Suction electrode. (b) Cross-sectional view of suction electrode in place, showing penetration of probe through epidermis. (c) Helical electrode, which is attached to fetal skin by corkscrew type action.

## Slide 25: Implantable Electrodes
Figure 5.15 Implantable electrodes for detecting biopotentials (a) Wire-loop electrode. (b) Silver-sphere cortical-surface potential electrode. (c) Multielement depth electrode.

## Slide 26: Microarray electordes
(c)
Tines
Base
Exposed tip
Contacts
Insulated leads
(b)
Base
Ag/AgCl electrodes
Ag/AgCl electrodes
Base
Insulated leads
(a)
Contacts
Figure 5.16 Examples of microfabricated electrode arrays. (a) One-dimensional plunge electrode array (after Mastrototaro et al., 1992), (b) Two-dimensional array, and (c) Three-dimensional array (after Campbell et al., 1991).

## Slide 27: Metal Microelectrodes
Figure 5.17 The structure of a metal microelectrode for intracellular recordings.

## Slide 28: Microelectrodes
Figure 5.18 Structures of two supported metal microelectrodes (a) Metal-filled glass micropipet. (b) Glass micropipet or probe, coated with metal film.

## Slide 29: Glass Micropipet Electrodes
Figure 5.19 A glass micropipet electrode filled with an electrolytic solution (a) Section of fine-bore glass capillary. (b) Capillary narrowed through heating and stretching. (c) Final structure of glass-pipet microelectrode.

## Slide 30: Yarı-iletken Mikro-elektrodes
Bonding pads
Si substrate
Exposed tips
Lead via
Channels
Electrode
Silicon probe
Silicon chip
Miniature
insulating
chamber
Contact
metal film
Hole
SiO2 insulated
Au probes
Silicon probe
Exposed
electrodes
Insulated
lead vias
(b)
(d)
(a)
(c)
Figure 5.20 Different types of microelectrodes fabricated using microelectronic technology (a) Beam-lead multiple electrode. (Based on Figure 7 in K. D. Wise, J.B. Angell, and A. Starr, "An Integrated Circuit Approach to Extracellular Microelectrodes." Reprinted with permission from IEEE Trans. Biomed. Eng., 1970, BME-17, pp. 238-246. Copyright (C) 1970 by the institute of Electrical and Electronics Engineers.) (b) Multielectrode silicon probe after Drake et al. (c) Multiple-chamber electrode after Prohaska et al. (d) Peripheral-nerve electrode based on the design of Edell.

## Slide 31: Eşdeğer devre
Metal rod
Tissue fluid
Membrane
potential
N
C
B
B
(a)
N = Nucleus
C = Cytoplasm
(b)
Emp
Membrane
and
action
potential
Cma
Rma
Cd + Cw
Ema - Emb
E
0
A
(c)
B
A
A
Reference
electrode
Rmb
Rma
Emb
Ema
Emp
Ri
Re
Cmb
Cma
Cdi
Cd2
Rs
Cw
Insulation
Cd
Cell
membrane
+
+
+
+
+
+
+
+
+
+
+
+
+
+
+
+
+
+
+
+
+
+
-
-
-
-
-
-
-
-
-
-
-
-
-
-
-
-
-
-
-
-
-
-
Figure 5.21 Equivalent circuit of metal microelectrode (a) Electrode with tip placed within a cell, showing origin of distributed capacitance. (b) Equivalent circuit for the situation in (a). (c) Simplified equivalent circuit. (From L. A. Geddes, Electrodes and the Measurement of Bioelectric Events, Wiley-Interscience, 1972. Used with permission of John Wiley and Sons, New York.)

## Slide 32: Eşdeğer devre
Ema
Rma
Rt
Ri
Re
(b)
Emp
Emb
Rmb
Cmb
Ej
Et
Cma
Cd
A
B
Rt
Em
A
B
Membrane
and
action
potential
(c)
Emp
Em = Ej + Et + Ema- Emb
Cd = Ct
0
Cell
membrane
Tip
+
-
+
-
+
+
+
-
-
+
-
-
+
-
+
-
+
+
+
+
+
-
-
-
-
-
+
+
+
-
-
-
+
-
-
+
+
+
+
-
-
-
Taper
Internal electrode
Glass
A
B
To amplifier
Electrolyte
in
micropipet
Stem
(a)
Reference
electrode
Cell membrane
Cytoplasm
N = Nucleus
N
Environmental
fluid
Cd
Figure 5.22 Equivalent circuit of glass micropipet microelectrode (a) Electrode with its tip placed within a cell, showing the origin of distributed capacitance. (b) Equivalent circuit for the situation in (a). (c) Simplified equivalent circuit. (From L. A. Geddes, Electrodes and the Measurement of Bioelectric Events, Wiley-Interscience, 1972. Used with permission of John Wiley and Sons, New York.)

## Slide 33: Gerilim ve Akım Uyarıları
(a)
Polarization
potential
Polarization
Polarization
u
u
i
i
t
t
t
t
Ohmic
potential
(b)
Polarization
potential
Figure 5.23 Current and voltage waveforms seen with electrodes used for electric stimulation (a) Constant-current stimulation. (b) Constant-voltage stimulation.