## Slide 1: Biyomedikal Algılayıcılar (Sensors), Dönüştürücüler (Transducers) ve Biyopotensiyellerin Algılanması
Çalışma Prensipleri ve Uygulamaları
Yalova Üniversitesi, 1 Kasım 2013

## Slide 2: Kaynaklar
http://www.unc.edu/~finley/BME422/Webster/c01.pdf
http://www.unc.edu/~finley/BME422/Webster/c02.pdf
http://megep.meb.gov.tr/mte_program_modul/modul_pdf/522EE0153.pdf

## Slide 3: 1 Kasım 2013 Gününün Konuları
İşaret işlemeyi etkileyen ölçüm sistemi parametreleri
Değişken direnç temelli algılayıcılar
Endüktif algılayıcılar
Kapasitif ve piezo elektrik algılayıcılar
Isıl çift (Thermocouple)
Alan etkili algılayıcı
Optik dönüştürücüler

## Slide 4: Basitleştirilmiş bir ölçüm sistemi
Ölçülen
Algılayıcı
Veri haberleşmesi
Veri göstergeleri
Etkileyici
İşaret şekillendir
İşaret işleme
Veri depola
Çıktılar
Tıbbi cihazlar genel olarak vücuttan alınan işaretleri izlerler. Bu işaretler analog olup önce elektriksel işaretlere çevrilir, işlenir ve bilgisayarla analiz edebilmek için sayısala çevrilir.

## Slide 5: Algılayıcılar
Algılayıcı tipleri
  Elektrotlar: elektriksel olan işaretler için
  Dönüştürücüler (transducers): kuvvet, basınç, sıcaklık gibi elektriksel olmayanları elektriksel işaretlere çeviren
Aktif ve pasif
  Aktif algılayıcı: Dışarıdan AC veya DC bir elektriksel güce ihtiyaç duyar; termistör gibi
  Pasif algılayıcı: kendi enerjisini sağlar veya ölçtüğü kaynaktan alır; ısıl çift (temokapl) gibi

## Slide 6: Algılayıcı Hata Kaynakları
Hata: ölçülen değerle gerçek değer arasındaki fark
5 tip hata kaynağı
  Yerleştirme hatası: algılayıcının yerleştirilmesi anında oluşur
  Uygulama hatası: kullanıcı hataları
  Karakteristik (özellik) hatası: cihazın bağıl hataları
  Dinamik hata: çoğu cihazlar statik konumda kalibre edilmiştir. Bir termistörde değer okuyacaksanız onun ortamla aynı sıcaklığa gelmesini beklemelisiniz
  Çevresel hata: sıcaklık, nem gibi çevresel etkenlerden oluşan hatalar

## Slide 7: Ölçü Aletlerinin Statik Özellikleri
Önemli ölçüm terimleri:
Doğruluk (Accuracy): Gerçek değere yakınlık
Hata (Error): gerçek değerden sapma
Hassasiyet (Precision): Ayırt edilebilen rakamlar (2,434 V 2,43 V dan daha hassas)
Çözünülürlük (Resolution): cevap verilen (ayırt edilebilen) en küçük değişim
Tekrarlanabilirlik (Repeatability): tekrarlanabilirliğin bir ölçüsü
İstatistiksel kontrol
Duyarlılık (Sensitivity): çıkış/giriş

## Slide 8: Duyarlılık (Sensitivity)
(a)
(b)
Duyarlılık (sensitivity) = çıkış / giriş
A low-sensitivity sensor has low gain.
A high sensitivity sensor has high gain

## Slide 9: Örnek
Input
Output
Input
Output
Hangisi daha duyarlı? Sol taraftakinde x deki degişikliğe karşı y deki daha fazla olduğundan o daha duyarlıdır.

## Slide 10: Algılayıcı (Sensor) Terminolojisi
Duyarlılık Hatası = karakteristik eyim eyrisinden uzaklaşma
Ideal Curve
Sensitivity Error
Output
Input

## Slide 11: Sürüklenmeler
Intercept b
Dxd
D
y
D x'd
D y'
y (Output)
y = mxd + b
xd (Input)
(a)
Slope m =
Dy
Dxd
Total error due to drift
Characteristic with zero and sensitivity drift
+ Zero
drift
+ Sensitivity
drift
- Sensitivity drift
- Zero drift
(b)
y (Output)
xd (Input)
Duyarlılık (sensitivity) = çıkış / giriş
Sıfır ve Duyarlılık Sürüklenmeleri (Zero and Sensitivity drifts)

## Slide 12: Dengeleme (Offset) Hatası
Dengeleme (Offset) hatası = Çıkış sıfır olması gerektiğinde olan değer
  Karakteristik eyrinin eğimi değişmez ama y eksenini kesim noktası merkezden uzaktadır.
Zero offset error
Offset Error
Input
Input
Output
Output

## Slide 13: Sıfır Dengeli ve Dengesiz İşaretler
(b) An input signal with dc offset.
(a) An input signal without dc offset

## Slide 14: Karışan ve Değiştiren Girişlerden Etkilenen İşaretler
(a)
(b)
(c)
(a) Original waveform. (b) An interfering input may shift the baseline. (c) A modifying input may change the gain.

## Slide 15: Doğrusallık (linearity)
xd (Input)
B% of full scale
A% of reading
Overall tolerance band
Least-squares
straight line
(a)
(b)
Point at which
A% of reading = B% of full scale
y (Output)
x1
(x1 + x2)
y1
x2
Kx1
Ky1
y2
Linear
system
Linear
system
Linear
system
Linear
system
and
and
(y1 + y2)
Doğrusallık (Linearity) = Ölçüm veya kalibrasyon sonucu elde edilen eyrinin ideal eyriden sapma miktarı.

## Slide 16: Doğrusal ve Doğrusal Olmayan Sistemler
(a)
(b)
(a) A linear system fits the equation y = mx + b. Note that all variables are italic. (b) A nonlinear system does not fit a straight line.

## Slide 17: Linearity
Nonlinearity (%) = (Din(Max) / INfs) * 100%
  Nonlinearity is percentage of nonlinear
  Din(max) = maximum input deviation
  INfs = maximum full-scale input
Input
Output
Ideal
Measure
Din(Max)
Full Scale Input

## Slide 18: Doğrusallığın Belirlenmesi
(a)
(b)
The one-point calibration may miss nonlinearity.
The two-point calibration may also miss nonlinearity.

## Slide 19: Dynamic Linearity
Measure of a sensor’s ability to follow rapid changes in the input parameters. Difference between solid and dashed curves is the non- linearity as depicted by the higher order x terms
F(x) = mx + K
F(x)* = ax + bx2+cx4+ . . . +K
Input X
Output
F(x)
K
F(x) = mx + K
F(x)* = ax + bx3+cx5+ . . . +K
Input X
Output
F(x)
K

## Slide 20: Dynamic Linearity
Asymmetric = F(x) != |F(-x)| where F(x)* is asymmetric around linear curve F(x) then
  F(x) = ax + bx2+cx4+ . . . +K offsetting for K or you could assume K = 0
Symmetrical = F(x) = |F(-x)| where F(x) * is symmetric around linear curve F(x) then
  F(x) = ax +bx3 + cx5 +. . . + K offsetting for K or you could assume K =0

## Slide 21: Hysteresis
Hysteresis = measurement of how sensor changes with input parameter based on direction of change

## Slide 22: Histeriz Etkisi
A hysteresis loop. The output curve obtained when increasing the measurand is different from the output obtained when decreasing the measurand.

## Slide 23: Hysteresis
The value B can be represented by 2 values of F(x), F1 and F2. If you are at point P then you reach B by the value F2. If you are at point Q then you reach B by value of F1.
Input = x
Output = F(x)
B
F1
F2
P
Q

## Slide 24: Giriş Büyüklüğü
Ölçülebilecek en küçük değer
Doğrusal olarak ölçülebilecek en büyük değer
Dayanılabilecek en yüksek değer

## Slide 25: Sensor Terminology
Range = Maximum and Minimum values of applied parameter that can be measured.
  If an instrument can read up to 200 mmHg and the actual reading is 250 mmHg then you have exceeded the range of the instrument.

## Slide 26: Sensor Terminology
Dynamic Range: total range of sensor for minimum to maximum. ie if your instrument can measure from -10V to +10 V your dynamic range is 20V
Precision = Degree of reproducibility denoted as the range of one standard deviation σ
Resolution = smallest detectable incremental change of input parameter that can be detected

## Slide 27: Girişin Sınırları
(a) An input signal which exceeds the dynamic range.
(b) The resulting amplified signal is saturated at 1 V.

## Slide 28: Accuracy
Accuracy = maximum difference that will exist between the actual value and the indicated value of the sensor
Xo
Xi

## Slide 29
(a)
(b)
Figure 1.16 Data points with (a) low precision and (b) high precision.

## Slide 30
Figure 1.17 Data points with (a) low accuracy and (b) high accuracy.
(a)
(b)

## Slide 31
Figure 1.19 For the normal distribution, 68% of the data lies within ±1 standard deviation. By measuring samples and averaging, we obtain the estimated mean, which has a smaller standard deviation sx.  is the tail probability that xs does not differ from  by more than .

## Slide 32
Figure 1.20 A typical Poisson distribution for m = 3.

## Slide 33
Conclusion | Real situation | 
 | H0 true | Ha true
Accept H0 | Correct decision | Type II error, p = b
Reject H0 | Type I error, p = a | Correct decision
Table 1.8 The four outcomes of hypothesis testing.

## Slide 34
Test result | Has condition? | 
 | No | Yes
Negative | True negative (TN) | False negative (FN)
Postitive | False positive (FP) | True positive (TP)
Table 1.9 Equivalent table of Table 1.8 for results relating to a condition or disease.

## Slide 35
Figure 1.21 The test result threshold is set to minimize false positives and false negatives.

## Slide 36: Giriş Empedansı
Xd1 – baskı (efor) değişkeni
Xd2 – akış (flow) değişkeni
Xd1 x Xd2 – güç boyutunda
Zx = Xd1 (baskı) / Xd2 (akış) – giriş empedansı
Yx = 1 / Zx – giriş admitansı
Sistem
Giriş
Çıkış
Efor
Akış

## Slide 37: Frequency Response of Ideal and Practical System
When you look at the frequency response of an instrument, ideally you want a wideband flat frequency response.
Frequency (w) radians per second
Av
Av = Vo/Vi
1.0

## Slide 38: Frequency Response of Ideal and Practical System
In practice, you have attenuation of lower and higher frequencies
FL and FH are known as the –3 dB points in voltage systems.
Av
Av = Vo/Vi
1.0
0.707
FL
FH
Frequency (w) radians per second

## Slide 39: Examples of Filters
Ideal Filter has sharp cutoffs and a flat pass band
Most filters attenuate upper and lower frequencies
Other filters attenuate upper and lower frequencies and are not flat in the pass band

## Slide 40: Bir Elektrokardiografın Frekans Cevabı

## Slide 41: Sıfır Derece Sistem
Doğrusal bir potansiyometre, sıfır dereceden bir sisteme bir örnektir.
Bu sistemin Lineer statik karakteristiği.
Basamak yanıtı girişle orantılıdır.
Sinüzoidal frekans yanıtı sıfır faz kayması ile sabittir.

## Slide 42: Birinci Derece Sistem
t
1
(c)
(a)
C
+
-
+
-
y(t)
Output y(t)
Input x(t)
Slope =
K
= 1
(b)
Y (jw)
X (jw)
Log
scale
1.0
0.707
Log scale w
(d)
0°
- 45°
-90°
Log scale w
t
1
0.63
L
S
L
S
wS
wL
f
L
S
x(t)
x(t)
y(t)
R
A low-pass RC filter, an example of a first-order instrument.
Static sensitivity for constant inputs.
Step response for larger time constants (L) and small time constants (S).
Sinusoidal frequency response for large and small time constants.

## Slide 43: İkinci Derece Sistem
Output
displacement
Output
y
(
t
)
(b)
(a)
(d)
(c)
1
Ks
x(t)
y(t)
yn
yn + 1
Resonance
2
Log
scale
1
2
-90°
0.5
1
2
-180°
1
0.5
0.5
Log scale w
Log scale w
K
1
t
t
Input x(t)
Slope K =
1
Ks
Input
Force x(t)
0
0°
wn
wn
Y (jw)
X (jw)
f
y(t)
Force-measuring spring scale, an example of a 2nd order instrument.
Static sensitivity.
Step response for
  overdamped (aşırı sönümlü) case  = 2,
  critically damped (kritik sönümlü) case  = 1,
  underdamped (zayıf sönümlü) case  = 0.5.
Sinusoidal steady-state frequency response,  = 2,  = 1,  = 0.5.

## Slide 44: Yaygın tıbbi ölçenler
Measurement | Range | Frequency, Hz | Method
Blood flow | 1 to 300 mL/s | 0 to 20 | Electromagnetic or ultrasonic
Blood pressure | 0 to 400 mmHg | 0 to 50 | Cuff or strain gage
Cardiac output | 4 to 25 L/min | 0 to 20 | Fick, dye dilution
Electrocardiography | 0.5 to 4 mV | 0.05 to 150 | Skin electrodes
Electroencephalography | 5 to 300  V | 0.5 to 150 | Scalp electrodes
Electromyography | 0.1 to 5 mV | 0 to 10000 | Needle electrodes
Electroretinography | 0 to 900  V | 0 to 50 | Contact lens electrodes
pH | 3 to 13 pH units | 0 to 1 | pH electrode
pCO2 | 40 to 100 mmHg | 0 to 2 | pCO2 electrode
pO2 | 30 to 100 mmHg | 0 to 2 | pO2 electrode
Pneumotachography | 0 to 600 L/min | 0 to 40 | Pneumotachometer
Respiratory rate | 2 to 50 breaths/min | 0.1 to 10 | Impedance
Temperature | 32 to 40 °C | 0 to 0.1 | Thermistor

## Slide 45: Kan Basıncı Algılayıcısının Özellikleri
Specification | Value
Pressure range | –30 to +300 mmHg
Overpressure without damage | –400 to +4000 mmHg
Maximum unbalance | ±75 mmHg
Linearity and hysteresis | ± 2% of reading or ± 1 mmHg
Risk current at 120 V | 10 A
Defibrillator withstand | 360 J into 50 
Kan basıncı algılayıcısının özellikleri akademi, sanayi, hastaneler ve hükümetten bireylerden oluşan bir komite tarafından belirlenmektedir.

## Slide 46: Bir Elektrokardiografın Özellikleri
Specification | Value
Input signal dynamic range | ±5 mV
Dc offset voltage | ±300 mV
Slew rate | 320 mV/s
Frequency response | 0.05 to 150 Hz
Input impedance at 10 Hz | 2.5 M
Dc lead current | 0.1 A
Return time after lead switch | 1 s
Overload voltage without damage | 5000 V
Risk current at 120 V | 10 A
Elektrokardiyografın teknik özellikleri bir komite tarafından kabul edilmektedir.

## Slide 47: İşaret ve enerji dönüşüm aygıtlarını gösteren 6x6 matris
Giriş→ Çıkış↓ | Işınım | Mekanik | Isıl | Elektrik | Manyetik | Kimyasal
Işınım | Optik filtre | Golay hücre det. | Solar metre | Güneş hücresi |  | Foto grafik işlem
Mekanik | Çakmak taşı | Dişli kutusu |  | Elektrik jen. |  | 
Isıl | Duyarlı LCD | Bimetal | Isı değişimi | Isıl çift |  | 
Elektrik | LED | Hoparlör | Soğutucu eleman | Mosfet | Sargı | Batarya
Manyetik | Manyete optik mod. | Manyetik debriyaj | Adıabatik anti manyet. | Hall levhası | Manyetik devre | 
Kimyasal | Kandil | Patlarlı motor | Gaz sobası | PH ölçüm hücresi |  | Kimyasal işlem

## Slide 48: Değişken Direnç Temelli Algılayıcılar / Dönüştürücüler
Potansiyometreli yer değiştirme
Gerilme ölçerler (strain gauge)
Termorezistif
PTC ve NTC termistörler
Foto direnç
Elektriksel iletkenlik (direnç)

## Slide 49: Potensiyometreler
Reosta

## Slide 50: Yer Değiştirme Ölçümlerinde Kullanılan Üç Tip Potensiyometre
Doğrusal
Tek turlu
Çok turlu
Helix-helezon

## Slide 51: Weston Köprüsü

## Slide 52: Devre Konfigürasyonu
E0 = VAC
VA = EbxR4/(R1+R4)
VC = EbxR3/(R2+R3)
E0 = VAC = VA – VC =

## Slide 53: Wheatstone Bridge
Basic Wheatstone Bridge uses one resistor in each of four arms where battery excites the bridge connected across 2 opposite resistor junctions (A and B). The bridge output Eo appears across C and D junction.
Es
R1
R2
R4
R3
EC
ED
+
-
A
B
Eo
Eo
R1
R3
R2
R4
EC
ED
Es
+
-
+
-

## Slide 54: Sıfırlama Modunda Çalışma
At balance:
R2R4 = R1R3 or R1/R4 = R2/R3 and the output voltage is zero

## Slide 55: Sapma Modunda Çalışma
All resistors can very around their nominal values as R1 + R1, R2 + R2, R3 + R3 and R4 + R4. Sensitivity of the output voltage to either one of the resistances can be found using the sensitivity analysis as follows

## Slide 56: Eşdeğer Devre
ETh = E0 = VAC (open circuit)
RTh = R1//R4 + R2//R3
Ig = E0/(RTh + Rg)
Eg = E0Rg/(RTh + Rg)
In case of open-circuit (Rg) Eg = E0

## Slide 57: Resistif Gerilme Ölçerler
Katlanmamış gerilim ölçer
(strain gage)
Katlanmış gerilim ölçer
(strain gage)

## Slide 58: Strain Gauges
Compression (Sıkışma) = decrease in length by DL and an increase in cross sectional area.
Rest Condition
L = length
L - DL = length
Compression

## Slide 59: Strain Gauges
Tension (Gerilme) = increase in length by DL and a decrease in cross section area.
Rest Condition
L = length
Tension
L + DL = length

## Slide 60: Resistance of a metallic bar is given in length and area
  where
    R = Resistance units = ohms (Ώ)
    ρ = resistivity constant unique to type of material used in bar units = ohm meter (Ώm)
    L = length in meters (m)
    A = Cross sectional area in meters2 (m2 )

## Slide 61: Resistance of a metallic bar is given in length and area
Example: find the resistance of a copper bar that has a cross sectional area of 0.5 mm2 and a length = 250 mm note the resistivity of copper is 1.7 x 10-8Ώm

## Slide 62: Piezoresistivity
Piezoresistivity = change in resistance for a given change in size and shape denoted as h
Resistance in tension =
Resistance increases in tension
    L = length; ΔL = change in L; ρ = resistivity
    A = Area; ΔA = change in A

## Slide 63
Resistance in compression =
Resistance decreases in compression
    L = length; ΔL = change in L; ρ = resistivity
    A = Area; ΔA = change in A

## Slide 64: Example of Piezoresistivity
Thin wire has a length of 30 mm and a cross sectional area of 0.01 mm2 and a resistance of 1.5Ώ.
A force is applied to the wire that increases the length by 10 mm and decreases cross sectional area by 0.0027 mm2
Find the change in resistance h.
  Note: ρ = resistivity = 5 x 10-7 Ώm

## Slide 65: Example of Piezoresistivity

## Slide 66: Example of Piezoresistivity
Note: Change in Resistance will be approximately linear for small changes in L as long as ΔL<<L.
If a force is applied where the modulus of elasticity is exceeded then the wire can become permanently damaged and then it is no longer a transducer.

## Slide 67: Gauge Factor
Gauge Factor (GF) = a method of comparing one transducer to a similar transducer

## Slide 68: Gauge Factor
where
  GF = Gauge Factor unitless
  ΔR = change in resistance ohms (Ώ)
  R = unstrained resistance ohms (Ώ)
  ΔL = change in length meters (m)
  L = unstrained length meters (m)

## Slide 69: Gauge Factor
    Where ε strain which is unitless
GF gives relative sensitivity of a strain gauge where the greater the change in resistance per unit length the greater the sensitivity of element and the greater the gauge factor.

## Slide 70: Example of Gauge Factor
Have a 20 mm length of wire used as a string gauge and has a resistance of 150 Ώ.
When a force is applied in tension the resistance changes by 2Ώ and the length changes by 0.07 mm.
Find the gauge factor:

## Slide 71: Katlanmamış gerilim ölçerler

## Slide 72: Basınç algılayıcıları için bir Weston köprüsü

## Slide 73: Katlanmış gerilme ölçerler
A bonded gage
Fixing the gage

## Slide 74: Katlanmış ölçer örnekleri
Resistance-wire type
Foil type
Helical-wire type
K  2.0
R0 = 120  or 350 . 600  and 700  gages are also available

## Slide 75: Yarı-iletken gerilme ölçer birimleri
Unbonded, uniformly doped
Diffused p-type gage

## Slide 76
Integrated pressure sensor

## Slide 77: Typical Configurations
R2 = SG2
R4 = SG4
R3 = SG3
C
+
-
A
B
Vo
R1 = SG1
D
ES
Electrical Circuit
Mechanical Configuration
4 strain gauges (SG) in Wheatstone Bridge:

## Slide 78: Strain Gauge Example
Using the configuration in the previous slide where 4 strain gauges are placed in a wheatstone bridge where the bridge is balanced when no force is applied,
Assume a force is applied so that R1 and R4 are in tension and R2 and R3 are in compression.
Derive the equation to depict the change in voltage across the bridge and find the output voltage when each resistor is 200 Ώ, the change of resistance is 10 Ώ and the source voltage is 10 V
+

## Slide 79: Strain Gauge Example
Derivation:
R2 = R - h
R4 = R +h
R3= R-h
C
+
-
A
B
Eo
R1 = R +h
D
Circuit
Es
+
-

## Slide 80: Transducer Sensitivity (Dönüştürücü Duyarlılığı)
Transducer Sensitivity = rating that allows us to predict the output voltage from knowledge of the excitation voltage and the value of the applied stimulus units = μV/V*unit of applied stimulus
Example: if you have a force transducer calibrated in grams (unit of mass) which allows calibration of force transducer then sensitivity denoted as φ = μV/V*g (another ex φ = μV/V*mmHg)

## Slide 81: Transducer Sensitivity
To calculate Output Potential use the following equations:
  where
    Eo = output potential in Volts (V)
    E = excitation voltage
    φ = sensitivity μV/V*g
    F = applied force in grams (g)

## Slide 82: Transducer Sensitivity
Example: Transducer has a sensitivity of 10 μV/V*g, predict the output voltage for an applied force of 15 g and 5 V of excitation.

## Slide 83: Metal Dirençlerini Isıl Değişimi
Metallerde Sıcaklığa Karşı Direnç Değişimi

## Slide 84: Dirençsel Isı Ölçer Algılayıcılar
NTC çeşitleri

## Slide 85: RTD Kuvvetlendiricisi

## Slide 86: Köprü Tipi RTD Kuvvetlendiricisi

## Slide 87: Termistörler
0.1
1.0
10
100
0.10
1.0
Water
Air
0.1 mW
1 mW
10 mW
100 mW
1 W
100 W
1 kW
10 kW
10 kW
100 k W
1 M W
A
C
B
Current, mA
(b)
10.0
100.0
(a) Typical thermistor zero-power resistance ratio-temperature characteristics for various materials. (b) Thermistor voltage-versus-current characteristic for a thermistor in air and water. The diagonal lines with a positive slope give linear resistance values and show the degree of thermistor linearity at low currents. The intersection of the thermistor curves and the diagonal lines with the negative slope give the device power dissipation. Point A is the maximal current value for no appreciable self-heat. Point B is the peak voltage. Point C is the maximal safe continuous current in air.
- 50 0 50 100 150 200
0.001
0.01
0.1
1
10
100
1000
Temperature, ° C
(a)
Resistance ratio, R/R25º C

## Slide 88: Termistör Uygulama Devresi
If the temperature increases, the thermistor resistance decreases, yielding more current that flows through Rf, thus Vo increases.

## Slide 89: Fotodirenç (LDR) (Light Depending Resistors)
Fotodirencin (a)Yapısı, (b) Sembolü, (c) Kılıfı

## Slide 90: Elektriksel İletkenlik (Direnç) Algılayıcısı

## Slide 91: Hücre sayaçları: Coulter model F

## Slide 92: Bobinli Manyetik (Endüktif) Algılayıcılar
L = n2G, burada
n= bobinin tur sayısı
G = geometrik form faktörü
 = etkin geçirgenlik (effective permeability)

## Slide 93: Endüktif Algılayıcılar
L = n2G, where
n= number of turns of coil
G = geometric form factor
 = effective permeability
Self (öz) endüktans
Karşılıklı (Mutual) inductance
Farksal (Differential) transformer

## Slide 94: Bobinli mikrofonlar
Dinamik (bobinli, manyetik) mikrofonlar
Şeritli (bantlı) mikrofonlar

## Slide 95: Doğrusal farksal gerilim transformatörü (LVDT)

## Slide 96: LVDT

## Slide 97: Functional operation of a phase-sensitive demodulator (a) Switching function. (b) Switch. (c), (e), (g), (i) Several input voltages. (d), (f), (h), (j) Corresponding output voltages.

## Slide 98: For this phase sensitive demodulator with the switch down, the gain is +1 for the first half cycle set by the reference signal. For the second half cycle the op amp inverter yields a gain of -1.

## Slide 99: LVDT Kullanım Alanları

## Slide 100: LVDT Kullanım Alanları

## Slide 101: Kapasitif Algılayıcılar
+
Q
-
Q
x
Area =
A

## Slide 102: Kapasitif mikrofonlar
Kapasitif (kondansatör) mikrofonlar
Elektret (electret) mikrofonlar

## Slide 103
Capacitive sensor for measuring dynamic displacement changes

## Slide 104: Kapasitif ince film nem algılayıcısı

## Slide 105: Piezo elektrik algılayıcılar
Quartz kristalindeki piezo elektrik etki

## Slide 106: Piezo elektrik kristalli (kristal) mikrofonlar

## Slide 107
k is piezoelectric constant C/N
K is proportionality constant C/m
KS=K/C, V/m;  = RC, s

## Slide 108: Basamak yerdeğiştirmeye cevap

## Slide 109: Yüksek frekans cevabı
High-frequency circuit model for piezoelectric sensor. RS is the sensor leakage resistance and CS the capacitance. Lm, Cm and Rm represent the mechanical system.
Piezoelectric sensor frequency response.

## Slide 110: Piezoelektrik dönüştürücü ve yük (charge) kuvvetlendiricisi
R
+
FET
Piezo-electric
sensor
-
uo
C
is
isR
isC
dqs/ dt = is = K dx/dt
The charge amplifier transfers charge generated from a piezo-electric sensor to the op-amp feedback capacitor C.

## Slide 111: Figure 14.15 Isolation in a disposable blood-pressure sensor. Disposable blood pressure sensors are made of clear plastic so air bubbles are easily seen. Saline flows from an intravenous (IV) bag through the clear IV tubing and the sensor to the patient. This flushes blood out of the tip of the indwelling catheter to prevent clotting. A lever can open or close the flush valve. The silicon chip has a silicon diaphragm with a four-resistor Wheatstone bridge diffused into it. Its electrical connections are protected from the saline by a compliant silicone elastomer gel, which also provides electrical isolation. This prevents electric shock from the sensor to the patient and prevents destructive currents during defibrillation from the patient to the silicon chip.
Gel
Clear plastic
4 cm
Saline
Flush valve
IV tubing
Electrical cable
Silicon chip
To patient

## Slide 112: Isıl Çift (Thermocouple)
(a) çalışma prensibi (b) Yapısı

## Slide 113: Thermocouple çeşitleri

## Slide 114: Thermocouple devreleri
(a) Peltier emf. (b) Law of homogeneous circuits. (c) Law of intermediate metals. (d) Law of intermediate temperatures

## Slide 115: Thermocouple’la ölçme
The hot junction is at the thermocouple. The LT1025 electronic cold junction compensates for ambient temperature changes. The noninverting amplifier provides a high input impedance and high gain.

## Slide 116: Alan Etkili Algılayıcı

## Slide 117: Optik Dönüştürücüler
Fotodirenç
Fotodiyot
Işık yayan diyot (light emitting diode – led)

## Slide 118: Fotodiyot: Karakteristik, sembol, yapı ve kılıfı

## Slide 119: Fotodiyodun kullanım alanları

## Slide 120: LED’in Yapısı, Sembolü, ve Kılıfı

## Slide 121: İnsan gözünün algılayabildiği dalga boyu sınırları

## Slide 122: İki ve üç renkli diyotların yapısı

## Slide 123: İnfrared (kızılötesi) diyodun kullanım alanları

## Slide 124: Foto transistörler
(a) Yapısı, (b) Sembolleri, (c) Dış görünüşü

## Slide 125: (a) 300 K deki bir cismin yaydığı ışığın sol dikey eksende şiddetinin tayfının, sağ dikey eksende toplam enerjinin yüzdesinin dalga boyuna göre değişimi. (b) Bir dizi optik malzemenin iletim spektrumu. (c) Foton ve termal dedektörler spektral duyarlılığı.
1
2
3
Wavelength, mm
1
0
10
50
100
10
5
0.001
0.002
0.003
0.00312
10
(a)
(b)
(c)
Wavelength, mm
15
20
T = 300 K
m= 9.66 mm
25
20
40
60
80
100%
Fused silica
Sapphire
Arsenic trisulfide
Thallium
bromide
iodine
Indium antimonide (InSb)
(photovoltaic)
Lead sulfide (PbS)
All thermal detectors
Wavelength, mm
100
0
20
60
100
4
5
6
7
8
% Total power
Spectral radient emittance, W-cm-2·mm-1

## Slide 126
Figure 2.17 Stationary chopped-beam radiation thermometer

## Slide 127
Figure 2.18 The infrared thermometer opens a shutter to expose the sensor to radiation from the tympanic membrane.

## Slide 128: Figure 2.19 Details of the fiber/sensor arrangement for the GaAs semiconductor temperature probe.

## Slide 129: Figure 2.20 (a) General block diagram of an optical instrument. (b) Highest efficiency is obtained by using an intense lamp, lenses to gather and focus the light on the sample in the cuvette, and a sensitive detector. (c) Solid-state lamps and detectors may simplify the system.

## Slide 130: Figure 2.21 Spectral characteristics of sources, filters, detectors, and combinations thereof (a) Light sources, Tungsten (W) at 3000 K has a broad spectral output. At 2000 K, output is lower at all wavelengths and peak output shifts to longer wavelengths. Light-emitting diodes yield a narrow spectral output with GaAs in the infrared, GaP in the red, and GaAsP in the green. Monochromatic outputs from common lasers are shown by dashed lines: Ar, 515 nm; HeNe, 633 nm; ruby, 693 nm; Nd, 1064 nm; CO2 (notshown), 10600 nm. (b) Filters. A Corning 5-65 glass filter passes a blue wavelength band. A Kodak 87 gelatin filter passes infrared and blocks visible wavelengths. Germanium lenses pass long wavelengths that cannot be passed by glass. Hemoglobin Hb and oxyhemoglobin HbO pass equally at 805 nm and have maximal difference at 660 nm. (c) Detectors. The S4 response is a typical phototube response. The eye has a relatively narrow response, with colors indicated by VBGYOR. CdS plus a filter has a response that closely matches that of the eye. Si p-n junctions are widely used. PbS is a sensitive infrared detector. InSb is useful in far infrared. Note: These are only relative responses. Peak responses of different detectors differ by 107. (d) Combination. Indicated curves from (a), (b), and (c) are multiplied at each wavelength to yield (d), which shows how well source, filter, and detector are matched. (e) Photon energy: If it is less than 1 eV, it is too weak to cause current flow in Si p-n junctions.

## Slide 131: Figure 2.22 Forward characteristics for p-n junctions. Ordinary silicon diodes have a band gap of 1.1 eV and are inefficient radiators in the near-infrared. GaAs has a band gap of 1.44 eV and radiates at 900 nm. GaP has a band gap of 2.26 eV and radiates at 700 nm.

## Slide 132: Figure 2.23 Fiber optics. The solid line shows refraction of rays that escape through the wall of the fiber. The dashed line shows total internal reflection within a fiber.
Air
n = 1.0
Coating
Fiber
n1
3
4
1
2
ic
n2

## Slide 133: Figure 2.24 Photomultiplier An incoming photon strikes the photocathode and liberates an electron. This electron is accelerated toward the first dynode, which is 100 V more positive than the cathode. The impact liberates several electrons by secondary emission. They are accelerated toward the second dynode, which is 100 V more positive than the first dynode, This electron multiplication continues until it reaches the anode, where currents of about 1 mA flow through RL.

## Slide 134: Figure 2.25 Voltage-current characteristics of irradiated silicon p-n junction. For 0 irradiance, both forward and reverse characteristics are normal. For 1 mW/cm2, open-circuit voltage is 500 mV, and short-circuit current is 0.8mA. For 10 mW/cm2, open-circuit voltage is 600 mV, and short-circuit current is 8mA.

## Slide 135
The photodiode yields a reverse current i proportional to light input. i must flow through Rf. Thus the op amp output Vo = iRf

## Slide 136: Gelecek Ders (22/11) İçin...
Bugünkü dersi detaylı olarak çalışınız ve anlaşılmayan noktaları belirleyiniz.
Megep sitesinden ‘Biyomedikal Algılayıcılar ve Dönüştürücüler’ notlarını indirip çalışınız.
http://www.unc.edu/~finley/BME422/Webster/c01.pdf ‘i indirip çalışınız.
Kısa bir sınava hazır olunuz.
Biyopotensiyel elektrotlar ve biyopotensiyellerin kaynakları konularını araştırıp çalışınız.